## Slide 1: Module 4: Improving the Quality and Efficiency of the Viral Load Testing Phase
DRAFT
Suitable for Laboratorians

## Slide 2: Module 4: Objectives
Compare and contrast the different viral load technologies commercially available
Operate an efficient viral load testing laboratory that produces timely and accurate viral load results
  Proper infrastructure and physical layout
  Optimal testing workflow
  Monitoring of key quality indicators
  Appropriate troubleshooting and corrective action when problems arise
  Timely result reporting
  Competent laboratory staff
  Complete laboratory documentation correctly

## Slide 3: Steps in the Testing Phase
Specimen receipt
Specimen check-in
Viral Load Testing
Result validation/ recording
Result
reporting
Testing Phase
Pre-testing Phase

## Slide 4
TAT 1
TAT 2
TAT 3
TAT 4
TAT 5
From sample pickup to sample delivery to testing laboratory
From time sample received at lab to result released to the clinic
  From receipt of results by clinic to return of results to clinician
  From receipt of results by clinician to return of results to patient for management of care
Keep Your Turn Around Time Short
Sample transport
Time spent inside testing lab
From clinic to clinician
From clinician to patient
From sample collection to sample pickup for transport
Wait time for sample pickup
Turn Around Time =
TAT 1 + TAT 2 + TAT 3 + TAT 4 + TAT 5

## Slide 5: Viral load testing technology

## Slide 6: Viral Load Testing
A quantitative test that measures the amount of HIV viral particles (viral load) in a patient’s blood
Highly sensitive molecular techniques including real-time PCR are utilized
VL is the earliest indicator that a patient is failing ART

## Slide 7: Technologies and Assays Available for Viral Load Testing
The four technologies available for VL testing are:
  Real-Time RT-PCR
  NASBA (Nucleic Acid Sequence Based Amplification)
  bDNA (Branched Chain DNA)
  Reverse Transcriptase Activity (none-molecular)
There are different assays and equipment from various manufacturers for each of the technologies listed above.
Notes: The NASBA technique differs from the PCR in that the amplification of nucleic acids takes place at a constant temperature (isothermal), usually at 410C . The RT assay differs from all the other VL techniques in that it is not a molecular-based technique but an ELISA assay that measures the activity of the reverse transcriptase. Because the RT is identical in all HIV subtypes, it is the only subtype independent VL assay.

## Slide 8: Commercially Available Viral Load Platforms
BioMérieux
Cavidi
Real-time
PCR
ELISA
Real-time
NASBA
NAT
Non-
NAT
Roche
Abbott
Siemens
Biocentric
Technology used
Automation
Throughput
Price
Infrastructure
Compatibility with DBS
Differences
Well validated
Sophisticated
Designed for plasma
Lab-based
Skilled staff
Similarities
Notes: NAT = Nucleic Acid Amplification Testing.
VL platforms share many similarities in terms of diagnostic performance, however, they differ in terms of type of technology, infrastructure requirements, automation, throughput and cost.

## Slide 9: Extraction
The 3 major viral load platforms (Roche, Abbott, and bioMérieux) use the same extraction method:
  During incubation of the lysed samples, all the target nucleic acid is captured by magnetic particles
  A magnetic device attracts all the magnetic particles, enabling the system to purify the nucleic acids through several washing steps
  The heating step releases the nucleic acids from the particles
  At the final step, the magnetic particles are separated from the eluate by the magnetic device

## Slide 10: Amplification
bioMérieux System | Real-Time NASBA
Roche Systems | Real-time RT-PCR Technology
Abbott m2000sp and m2000rt System | Real-time RT-PCR Technology

## Slide 11
Used to amplify RNA only
Primer incorporates T7 promoter and amplifies via T7 RNA polymerase
Uses isothermal conditions, removing time to cycle between temperatures
The NASBA reaction does not give false positives caused by genomic dsDNA, as in the case of RT-PCR
This makes it ideal for DBS testing
NASBA® Amplification
(bioMérieux System)
Notes: Initially, amplification causes an exponential rise in the number of amplicons as more RNA copies generated from each RNA target sequence, enter the amplification process. As the reaction progresses the primer pool will become depleted. After primer depletion, the accumulation of RNA transcripts proceeds solely due the transcription of double-stranded DNA, leading to a linear increase in amplicons. It is this phase of the reaction that is used for the determination of transcription products.

## Slide 12
bioMérieux System
NucliSens miniMAG/easyMAG and easyQ
Notes: The extraction process can be done with a manual version (miniMAG), which is suitable for low-volume testing (12 tests), or a semi-automated system (easyMAG) which is suitable for high-volume testing (24 tests/hour). The amplification process is done with the automated platform (easyQ) using real-time NASBA technology (48 tests/hour). In order to increase throughput two easyMAGs can be combined with one easyQ.

## Slide 13: BioMérieux Viral Load Assay NucliSENS EasyQ HIV-1 V2.0
Type of assay | real-time NASBA Isothermal Signal Amplification | 
Dynamic Range (copies/mL) | 25 – 10,000,000 | 
Specimen Type | Plasma (EDTA tube) and DBS (CE-marked) | 
Specimen volume | 100 to 1000μL | 
Area of HIV genome measured | GAG | 
HIV-1 subtypes measured | HIV-1 Group M subtypes A-J including CRF | 
Time for result | Estimated 1 hour (extraction) + 1.5 hours = 2,5 hours | 
Number of samples/ run (max.) | easyMAG: 24 (22 patient samples + 2 controls)1 EasyQ: 48 (44 patient samples + 4 controls) | 
Equipment required – RNA Extraction | miniMAG | easyMAG
Equipment required – Amplification/ Detection | EasyQ | 
1) Two controls per extraction batch of 24 samples
Notes: The NucliSENS EasyQ® HIV-1 v2.0 (Automated) has received WHO prequalification for its use with plasma. The DBS method is CE-marked.

## Slide 14: Real-time RT-PCR Amplification (Roche and Abbott)
Reverse Transcriptase (RT) is used to transcribe extracted RNA to cDNA, which is then amplified by PCR.
Detection is done via fluorescent oligonucleotide probes that bind specific target sequences in the HIV-1 RNA genome and a known quantity of quantitation standard (QS) RNA sequence.
The HIV-1 RNA signal is compared with the QS signal which allows for quantitation.
Notes: The COBAS® AmpliPrep/COBAS® TaqMan® HIV-1 Test, v2.0 has received WHO prequalification for its use with plasma samples.

## Slide 15
COBAS® Ampliprep/
COBAS®
TaqMan ® 96 Sample Format
Roche Systems
COBAS® Ampliprep/ COBAS®
TaqMan ® Analyzers
COBAS® Ampliprep/ COBAS®
TaqMan ® 48 Analyzer

## Slide 16: PCR: Roche Viral Load Assays COBAS® AmpliPrep/ COBAS® TaqMan® HIV-1
Type of assay | real-time PCR | 
Dynamic Range (copies/mL) | 20 – 10,000,000 | 
Specimen Type | Plasma (EDTA tube) and DBS (RUO) | 
Specimen volume | 1000 to 1050 μL (plasma) and one 70 mL spot (DBS) | 
Area of HIV genome measured | GAG, LTR | 
HIV-1 subtypes measured | HIV-1 Group M subtypes A-H, Group O, including CRF | 
Time for result (per run) | Estimated 2,5 hours + 3 hours = 5,5 hours | 
Number of samples/ run (max.) | TaqMan 48: 48 (42 patient samples + 6 controls)(1,2) TaqMan 96: 96 (84 patient samples + 12 controls)(1,2) | 
Equipment required – RNA Extraction | COBAS® AmpliPrep System | 
Equipment required – Amplification/ Detection | COBAS® TaqMan 48® Analyzer OR | COBAS® TaqMan 96® Analyzer
Three controls per run of up to 24 each (21 patient sample + 3 controls)
2 or 4 independent (respectively) 24-well thermocyclers
Notes: The COBAS® AmpliPrep/COBAS® TaqMan® HIV-1 Test, v2.0 is the latest version and has received WHO prequalification for its use with plasma samples. The use of DBS remains for Research use only (RUO).

## Slide 17
m2000sp
m2000rt
Abbott
M2000sp and m2000rt System

## Slide 18: PCR: Abbott Viral Load Assay RealTime HIV-1
Type of assay | real-time PCR | 
Dynamic Range (copies/mL) | 40 – 10,000,000 | 
Specimen Type | Plasma (EDTA tube) and DBS (RUO) | 
Specimen volume | 200 to 1000 μL (plasma) and one 70 mL spot (DBS) | 
Area of HIV genome measured | POL Integrase | 
HIV-1 subtypes measured | HIV-1 Group M subtypes A-H, Group O, Group N, including CRF | 
Time for result | Estimated 5 hours (extraction) + 2.5 hours (amplification) = 7.5 hours | 
Number of samples/run (max.) | m24sp: 24 (21 patient samples + 3 controls) (1) m2000sp: 96 (93 patient samples + 3 controls)(1) | 
Equipment required – RNA Extraction | m2000sp or | m24sp
Equipment required – Amplification/Detection | m2000rt | 
1) Three controls per batch or 96 deep well plate (21 or 93 patient sample + 3 controls)

## Slide 19: Summary of Commercially Available Platforms
Handout
4-1
From WHO Technical and Operational Considerations for Implementing HIV Viral Load Testing, July 2014.

## Slide 20: Selecting Viral Load Technologies for your Setting
The selection of viral load technologies should not rely solely on the technical specification and characteristics of a viral load platform
Consider specimen types and operational considerations:
    Efficiency of specimen transport system
    Expected testing volumes
    Human resources
    Technical support.
What should a country consider, besides costs, when selecting the viral load technologies and platforms?

## Slide 21: Knowledge Check
What is the molecular technique used for viral load testing by both Roche and Abbott? Select the most appropriate answer from below.
  NAT technology
  NASBA® Amplification
  Real-time RT-PCR
  bDNA
√
A
B
C
A and C
All

## Slide 22: Knowledge Check
What is the molecular technique used by NucliSens? Select the best answer.
  NASBA® Amplification
  bDNA
  Real-time RT-PCR
  All of the above
√

## Slide 23: viral load testing infrastructure and workflow

## Slide 24: Viral Load Testing Laboratory Infrastructure
2 – 3 dedicated rooms with minimal dust, preferably with temperature control (air-conditioner)
Sample Prep and extraction Room
  Sample prep extraction should be done preferably in a Biosafety Cabinet
MasterMix Room (can be reduced to a Clean-Air Box in ROOM 1 if space is limited)
  Master-Mix Preparation
  Only required if using a manual or semi-automated test method
Amplification and Detection Room (becomes highly contaminated and should be separate from ROOMS 1 and 2 in case of separate amplification and detection)
  PCR Amplification and Detection
  Results validation and documentation
To avoid contamination workflow in the laboratory MUST proceed in a unidirectional manner.
Sample Prep and extraction Room
MasterMix Room
Amplification and Detection Room
Notes: Useful WHO reference: Establishment of PCR laboratory in developing countries http://apps.searo.who.int/PDS_DOCS/B4745.pdf
PCR has a strong susceptibility to contamination, therefore, the most critical aspect in working in a PCR laboratory is to follow a unidirectional work flow. Wear different lab coats for each room. Never go back from Room 3 to Room 2 or 1 without changing coats.

## Slide 25: Viral Load Testing Laboratory Physical Set-up
bench space: Each room should have 3 – 4 meters bench space.
Storage: Reagent kits should be stored according to manufacture specified temperatures
Electricity:
  Steady current and dedicated power
  Uninterruptible power supply (UPS)
Biohazard waste disposal facilities: biohazard bags, bins, autoclave and incinerator
Notes: Useful WHO reference: Establishment of PCR laboratory in developing countries http://apps.searo.who.int/PDS_DOCS/B4745.pdf

## Slide 26: Suggested Laboratory Layout Sample Extraction and Master-Mix Prep. Rooms (1 & 2)
Master-Mix Preparation Area (ROOM 2)
This “clean area” can be a separate room or can consist of a “Closed Air Box” in one corner of the Sample Prep/Extraction Room (ROOM 1)
Sample Prep and Extraction Area
(ROOM 1)
Freezer
(Sample Storage)
80 cm
Sample Prep and Extraction
(preferably in a BSC)
200 cm
Closed Air
Box
90 cm
Fridge/freezer
(Reagent Storage)
80 cm
Bench for Centrifuge
70 cm
Bench area
Manual or Automated Extraction Instrument
Notes: Use negative pressure in the Sample Extraction Area (Room 1) to keep template nucleic acids in the room.
Use positive pressure in the Master-Mix Preparation Area (Room 2) to prevent the introduction of contamination.
Educational video: https://www.youtube.com/watch?v=V2JYy6-DE9c

## Slide 27: PCR: Suggested Laboratory Layout Sample Amplification and Detection Room (3)
Amplification and Detection Area
(ROOM 3)
100 cm
Result validation and documentationarea
200 cm
200 cm
Amplification/ detection instrument
Printer
Notes: The thermocycler /real-time PCR machine should be kept in this room. Gloves and laboratory coats should be worn at all times and removed before leaving the room to control amplicon contamination of other locations. All equipment used for amplification and product detection should be dedicated to this room, including adjustable pipettes with plugged, aerosol-barrier tips. This room should be kept under negative pressure to keep amplified nucleic acids in the room.

## Slide 28: Maximizing your Viral Load Testing Throughput
Scenario | Maximum Throughput | Your current level?
One Roche 96 machine, 8-hour shift, 2 people | 110 specimens | 
One Roche 96 machine, 12-hour shift, 2 people | 150 specimens | 
One Roche 96 machine, 24-hour shift, 2 people | 270 specimens | 
One Abbott m2000 machine, 8-hour shift, 2 people | 111 specimens | 
One Abbott m2000machine, 12-hour shift, 2 people | 180 specimens | 
One Abbott m2000machine, 24-hour shift, 2 people | 270 specimens | 
Brainstorm ways your can improve your throughput:
Are these number realistic to you?
What is the best way to achieve or even go beyond these numbers?

## Slide 29: Tips for Maximizing Workflow Efficiency
Use free downloadable app to generate barcodes for specimens’ lab ID to minimize transcription errors and data entry time
Use appropriate pipettes (2 ml instead of 1 ml) when transferring 1.1 ml plasma into the reaction tubes to save time and to avoid cross-contamination
Vortex plasma inside the BSC to minimize specimen travel
Use the same type of DBS cards to simplify procurement and training
Use perforated DBS cards instead non-perforated cards. Use disposable plastic pipettes to punch out the circles (instead of scissors)
Assign dedicated personnel to alert clinics of high viral load results on a daily basis
Minimize instrument down time between procedures using teams working in a manufacture assembly-line fashion.

## Slide 30: Knowledge Check
How should viral load testing work flow proceed? Select the best answer.
  From Extraction room to Master mix room to Amplification room
  From Master mix room to Extraction room to Amplification room
  From Amplification room to Extraction room to Master mix room
√

## Slide 31: Sample receipt, rejection, check-in and storage until testing

## Slide 32
Handout
4-2
3-10

## Slide 33
Handout
3-11

## Slide 34
Handout
4-3
3-5

## Slide 35: Preventing and Troubleshooting Problems

## Slide 36: Common Errors – How to Prevent them?
Errors may be caused by:
Improperly thawed/mixed/ dissolved reagent
Air bubbles in reagents and reagent lines
Sampling or reagent syringe problems
Pipette tips not fitting properly
A clogged pipette
Imprecise pipette
Unstable temperature and incubation
Variations in operator technique
Incorrect reconstitution of the control product
To minimize errors:
Strictly follow the SOP
Always perform instrument and laboratory equipment maintenance and calibrations
Ensure a consistent and appropriate testing environment
Ensure the competency of the testing personnel
Notes: Random errors are caused by unknown and unpredictable changes during testing.
These changes may occur in the measuring instruments or in the environmental conditions. Because random errors are unpredictable, singular events, they can be difficult or even impossible to detect. Random errors should only be considered if ALL alternative sources of error have been investigated.
Although random errors can be extremely hard to detect retrospectively, they can be avoided by preventative action and consideration of certain factors. During all testing events, be aware of and avoid the following:
This list is non-comprehensive. Random errors can occur at any moment during testing. Take every precaution to maintain consistency at every moment during testing.
To prevent random errors, avoid variation between testing events, always perform instrument and laboratory equipment maintenance and calibrations, and guarantee a consistent and appropriate testing
environment.

## Slide 37
Handout
4-4

## Slide 38: Activity 4A: Troubleshooting and Corrective Action
Purpose
To develop corrective action for each category of problem encountered during viral load testing.
What will you need?
Flipchart paper
Handout 4-4: When Errors Occur - Troubleshooting Guide
What will you do?
Divide into 7 groups - each group will be assigned one category of problem in the Troubleshooting Guide.
  Review your category of problem in the troubleshooting guide and discuss within your group the corrective action needed to address the problem
  Write your suggestions on a flipchart
  Select a spokesperson to present your corrective actions ( 2 minutes).
  Participate in large-group discussion
10 minutes
Handout
4-4
Notes: 38

## Slide 39: Activity 4A: Suggested Corrective Action
Handout
4-5

## Slide 40: Ensuring the quality of the testing phase

## Slide 41: Viral Load Laboratory Readiness Assessment Checklist
Handout
4-6

## Slide 42: Quality Indicators
Number of specimens rejected (by reason, by clinic)
Turn around time from specimen collection to specimen arrival at lab
Turn around time from specimen arrival to testing
Turn around time from testing to result return to clinic
Number of specimens tested
Number of invalid results (per week or month)
Number of Target Not Detected (TND)
Number of high viral load results (>1000 copies/ml)
Number of runs out of QC range
Number of days testing is halted (per month) due to instrument breakdown
Number of days testing is halted (per month) due to out-of-stock
Backlog – both number of specimens and length of time in storage

## Slide 43: Tracking Sample Rejection by Criteria and by Clinic Pinpoints the Source of the Problem
Reason for Rejection
Frequency of occurrences
Site name

## Slide 44: Ensuring the Competency of Testing Personnel
Test operators must be trained
Their competency must be assessed at least once a year
Handout
4-7

## Slide 45: Documentation and Records for Viral Load Testing
Handout
4-8

## Slide 46: Result recording and reporting

## Slide 47: Guidelines for Result Reporting (1)
Handout
4-9

## Slide 48: Guidelines for Result Reporting (2)
Handout
4-9
Results generated
Data clerk at lab reviews all approved results
  Record sample #, ART #, DOB, date of result and value, and facility name
  Data clerk prints the reports
  Email or call the POC at clinic with the information for entry into daily log book and high VL register
  Deliver results via email, SMS, fax, courier, etc.
  Verify results receipt at clinic
  Retain results in secure area
≥1000 copies/ml
<1000 copies/ml
  Verify receipt

## Slide 49: Key Messages
The selection of viral load technologies should not rely solely on the technical specification and characteristics of a viral load platform, but also specimen types and operational considerations, such as _________________, _______________, _______________, and technical support.
Although new viral load technologies are highly automated and sophisticated, strict adherence to ____, regular ___________, and _______________are crucial in guaranteeing accurate and reliable results.
Proper _____________, ____________, and _________ are critical aspects of a viral load laboratory in order to reduce the risk of contamination.
Testing personnel’s competency should be assessed at least _____________.

## Slide 50: Key Messages
The selection of viral load technologies should not rely solely on the technical specification and characteristics of a viral load platform, but also specimen types and operational considerations, such as efficiency of specimen transport system, expected testing volumes, human resources and technical support.
Although new viral load technologies are highly automated and sophisticated, strict adherence to SOPs, regular equipment maintenance, and personnel competency are crucial in guaranteeing accurate and reliable results.
Proper infrastructure, physical layout, and workflow are critical aspects of a viral load laboratory in order to reduce the risk of contamination.
Testing personnel’s competency should be assessed at least once a year.